## Slide 1: Multi-Altiz Alignment
Calibration Tool
One line

## Slide 2
70mm
360mm
40mm
40mm
40mm
60mm
60mm
15mm
60mm
60mm
Sketch 2D model one line
For information only. Dimensions are subject to change.

## Slide 3: Sketch 2D model
For information only. Dimensions are subject to change.

## Slide 4
3D model

## Slide 5: Material
The tool should have a matte finish to reduce reflections. Matte paint can be used:

## Slide 6: Constraints (1)
The Altiz must be ordered with the main camera at one end.
  Example:
Main
Secondary 1
Secondary 2
Secondary 3
Secondary 4
Secondary 3
Secondary 2
Secondary 1
Main
Secondary 4
Output Input
Output Input
Output Input
Output Input
Input Output
Input Output
Input Output
Input Output

## Slide 7: Constraints (1)
Secondary 2
Secondary 1
Secondary 3
Main
Secondary 4
Secondary 3
Main
Secondary 1
Secondary 4
Secondary 2
Output
Input
Output
Input
Output
Input
Output
Input
Output
Input
Output
Output
Input
Output
Input

## Slide 8: Constraints (2)
The Altiz should not have any rotation around their X and Z axis.

## Slide 9: Constraints (3)
Grab must only be performed on the Main camera.
Plugin support a maximum of 6 Altiz.

## Slide 10: Installation (1)
Required:
  MIL X 23H2 and U149 minimum .
  SP7 minimum.
  Image Processing module license.
To add a plugin in Capture Works:
  Download the plugin Zip folder on GitHub. (see next slide)
  Unzip. Inside you’ll find a PowerShell script (.ps1) and a folder containing the plugin.
  Open PowerShell as an administrator.
  Use the “cd” command to navigate to the directory containing your PowerShell script.
  Write .\MultiAltizInstallScript.ps1 and press enter. The script will check:
    Python installed in your machine.
    pip package install.
    Mil Python package. If not, it will install it.
  Once all checking are done it will copy the plugin folder in CaptureWorksPlugin folder.

## Slide 11: Installation (2)
Download the plugin Zip folder on GitHub
1- Click on “Code”:
2- Click on “Download ZIP”

## Slide 12: User guide
After opening CaptureWorks open the “Plugin manager” :
Select “multi_altiz” then clic on OK

## Slide 13: User guide
Allocate each secondaries camera and then allocate the main camera. You should see a new tab called “Multi AltiZ Plugin”:

## Slide 14: User guide
Extend the Multi AltiZ Plugin tab to see all the features it contains. (1)
Plugin’s name.
Sub tab allowing the user to set the plugin’s parameters.
Main camera name.
Control cameras parameters.
Sets the distance between the holes along Altiz’s X-axis.
Instruction and status of alignment.
Set whether you want to perform alignment.

## Slide 15: User guide
Extend the Multi AltiZ Plugin tab to see all the features it contains. (2)
Based on the IO, the plugin will automatically find topology and identifies the main camera (when it is allocated).

## Slide 16: User guide
Extend the Multi AltiZ Plugin tab to see all the features it contains. (3)
Sub tab allowing user to set plugin’s parameters.
Selects whether you want to save settings in a camera.
Note that when activated, it slows the plugin just before the grab. Make sure the “Acquisition…” status is on before moving the stage.
Keep the data of all the grabs and not only the merge.
Refresh values on the feature browser.
Setup the laser multiplexing mode: - Neighbour (default)
- Off
- Custom
Assign a unique color per PC grabbed by each AltiZ.

## Slide 17: User guide
Extend the Multi AltiZ Plugin tab to see all the features it contains. (4)
Set whether you want to perform alignment. When feature is set to “Alignment”, simply press the Single Grab button to performs the grab and alignment. Or set this feature to “Grab Only” when an alignment has already been made. The conveyor can be moved.

## Slide 18: User guide
Extend the Multi AltiZ Plugin tab to see all the features it contains. (5)
Indicates if the plugin is ready to align or not.

## Slide 19
Extend the Multi AltiZ Plugin tab to see all the features it contains. (6)
Set the distance between the holes along Altiz’s X-axis from the main camera.
40mm
40mm
40mm
100 mm
100 mm
Main
Secondary 1
Secondary 2

## Slide 20
Extend the Multi AltiZ Plugin tab to see all the features it contains. (6)
40mm
40mm
40mm
100 mm
100 mm
Secondary 2
Secondary 1
Main

## Slide 21
Extend the Multi AltiZ Plugin tab to see all the features it contains. (7)
Control cameras parameters.
Transformation values for the alignment. Automatically filled after an alignment but the user can enter his own values.
The transformation values are relative to the main camera frame.
Indicates position of the selected camera in the topology.
Indicates whether the selected camera is the main camera. If not, it indicates which camera it is connected to.
Example of a secondary camera.

## Slide 22: Setup time
Step 1: Position the cameras, making sure there's no pitch or yaw by using the setup view on CaptureWorks that will only shoot the laser.
Step 2: Place the calibration tool on the conveyor. The calibration tool doesn’t have to be perpendicular to direction of motion. One hole should be under each camera. If there is Roll around the Y-axis, it is ok if the camera sees more than one hole.

## Slide 23
Section of the calibration tool using Medium Altiz
Main Altiz
Secondaries

## Slide 24
Section of the calibration tool using Small Altiz
Main Altiz
Secondaries

## Slide 25: Runtime (Alignment ON)
Step 1: In the same instance of CaptureWorks, allocate all cameras, ending by the main camera, to obtain the complete topology on this camera.
Step 2: Set cameras parameters such as “Length World” or the “Motion Step World”.
Steps 3: On the main camera, extend “Multi Altiz Plugin” expander. There, set the “Distance” value.
Step 4: Set Operation to Alignment. Click on Single Grab to find the transformation matrices of each cameras.

## Slide 26: Runtime
Step 5: Start moving the conveyor.
Step 6: The merge point cloud is displayed (if View Mode is set to All cameras, then all grabs from each camera are display too). If the result is good, turn Operation to Grab Only. You can save the buffer for later processing. The transformation values are now saved on each camera.
Step 7: Place an object to scan and click on CaptureWorks Single Grab button. This will only perform the merge with the transformation values obtained during alignment.

## Slide 27: Runtime (Alignment OFF)
Step 1: In the same instance of CaptureWorks, allocate all cameras, ending by the main camera, to obtain the complete topology on this camera.
Step 2: Set cameras parameters such as “Length World” or the “Motion Step World”.
Step 3: Set Operation to Grab Only. Click on Single Grab. This will only perform the merge with the transformation values obtained during alignment.
Step 4: Start moving the conveyor.
Step 5: The merge point cloud is displayed (if View Mode is set to All cameras, then all grabs from each camera are display too). You can save the buffer for later processing.